## Slide 1: Flattening filter free photon beams Dosimetric characteristics and clinical implications
Dietmar Georg1,2
1Division Medical Radiation Physics, Department of Radiation Oncology
2Christian Doppler Laboratory for Medical Radiation Research for Radiation Oncology
Medical University Vienna & AKH Wien

## Slide 2: Outline of talk
Historical perspective
Dosimetric characteristics
Treatment planning aspects
  Dose calculation accuracy
  Clinical potential
Radiation protection
Radiobiology
Summary

## Slide 3: Vienna FFF history
Experimental FFF beam since 2005
  Replacement of scattering foil with 6mm Copper disk + “new beams” on console
  2 mm Stainless steel plate in 2010
  TPS commissioned as clinical unit
Clinical FFF unit in 2014
3

## Slide 4: Why using flattening filters?
Photon dose distribution produced by a target is strongly forward peaked
Clinically “useful” dose profiles are traditionally produced by flattening the raw beam with a so called flattening filter
Med Phys. 1991 (18) 519-521

## Slide 5: Flattening filter free treatment units
Kawachi et al MP 35 (2008)
Cyberknife
Tomotherapy
Jeraj et al MP 31 (2004)
Target exchanger
Purging magnet
Wedge filter
Ion Chamber
Helium atmosphere
Multileaf collimator (MLC)
Video camera
Filter revolver
Beam scanning magnets
  Scanditronix MM50 racetrack microtron
Head design of “conventional” C-arm based linacs is revisited

## Slide 6: Dosimetric properties of FFF beams
Increased dose rate
Off-axis spectral variation reduced
Small head scatter factor variation – IMRT with elongated/small segments
Contaminating electrons – smaller field size variation of dmax
Sixel et al MP (1994)
Kragl et al R&O 93 (2009)
Georg et al MP 37(2010)

## Slide 7: REMINDER - Multisource models in TPS
Multisource models:
  Focal source
  Extra focal source(s): primary collimator, FF scatter, collimator scatter
  Electron contamination
Chaney E L et al A Monte Carlo study of accelerator head scatter. Med. Phys. 21 (1994)
Distribution of origin site.
- clouds at target, primary collimator and flattening filter
Lateral distribution of origin sites, projected to a common distance from the target.
e-
De-
x
Olofsson, Georg, Karlsson R&O 67 (2003)

## Slide 8: Dose calculation in FFF mode
No statistically significant differences on-axis
Average γmean – off-axis | n | FF | FFF | p
γmean (MC) | 18 | 0.55 ± 0.24 | 0.45 ± 0.21 | 0.02
γmean (eCC) | 18 | 0.41 ± 0.24 | 0.36 ± 0.22 | 0.04
Kragl et al R&O 100 (2011)

## Slide 9: Clinical implications of FFF beam characteristics
What is the added value?
  3D-CRT
  S(B)RT
  IMRT / VMAT
How about treatment plan quality ?
Are there limitations ?
How about QA aspects ?
  Dosimetric issues
typical dose profile without FF

## Slide 10: Cranial SRT – FF vs. FFF
Treatment time | Uveal melanoma 5 x 10Gy @ 80% (10 beams) | Cerebral meta (<10ccm) 1 x 18Gy @ 80% (10 beams) | Cerebral meta (>10ccm) 2 x 10Gy @ 80% (10 beams) | Ependymom 25 x 2Gy @100% (9 beams) | Acoustic neurinoma 29 x 1.8Gy @100% (10 beams)
Monitor units | 1188 MU | 2096 MU | 1518 MU | 240 MU | 216 MU
FFF / FF Tot TT | 78% | 70% | 75% | 92% | 94%
Dose rate effect: Beam-on time reduced by ~ 50% BUT ….
Notes: all estimations for X6

## Slide 11: 4D effects: Inter- and Intra-fraction motion
11
11
Modified from Yan et al PMB (1997)

## Slide 12: SBRT with and without FFF-beams
IMRT FF
4903MU
VMAT FF
4977MU
CRT FF
2961MU
IMRT FFF
4730MU
VMAT FFF
3732MU
CRT FFF
2865MU
Clinical practice @ MedUni Wien: 15 Gy @ 65% isodose
Notes: 7-beam CRT [OMP]
6-IMRT (FF 19 Seg; FFF 15Seg) [Monaco]
2 partial arc VMAT (180° - 220°, 40° -180°) (FF 142 CP; FFF 119 CP) [Monaco]

## Slide 13: Real time motion tracking – paired kV – MV sets
    WIP: Offline evaluation on 20 SBRT patients to check feasibility for clinically used gantry angles
    WIP: DIBH Monitoring by means of imaging
Furtado et al Acta Oncol (2013)

## Slide 14: Treatment plan quality of “FFF plans”
  Pareto optimality:
    The improvement of an objective can only be achieved by deteriorating another objective
    Can be used for the comparison of different modalities
    Is a desirable property of treatment plans
speed
power
weak
strong
slow
fast
courtesy Ralf Dreindl
fast but weak
slow but strong
neither fast nor strong
Pareto Optimal Front
  Classical treatment planning studies depend a big deal on experience …

## Slide 15: Pareto optimal fronts – Prostate
Symbols depict treatment plans
Lechner et al R&O 109 (2013)
Notes: 50/60 Gy in 28 Fractions
Initial plans very similar

## Slide 16: Pareto optimal fronts – Head & Neck case
Symbols depict treatment plans
Lechner et al R&O 109 (2013)
Notes: 50/60 Gy in 28 Fractions
Initial plans very similar

## Slide 17: IMRT Pre-treatment verification
Kragl et al. R&O 100 (2011)
Lang et al. MP 39 (2012)
Diode arrays : No DR effects reported for high dose rates
2D arrays were shown to suitable up to 24 Gy/min, fractional doses 1.8 to 24 Gy
„2nd generation“ arrays revised for FFF
Notes: Mention ArcCheck and Delta4

## Slide 18: Radiobiologic considerations….
Sorensen et al R&O 101 (2011)
“…..This suggests that the high instantaneous dose rates of FFF linear accelerators will not influence either treatment outcome nor treatment toxicity.”
“….. More data is needed to better understand the effect of modified dose rate on tumour as well as on normal tissue …….”
Lohse et al R&O 101 (2011)
Dose rate and PRF !!!
Realistic oxygen conditions?

## Slide 19: Radiation Protection aspects
We know for a long time that patients undergoing radiation therapy are exposed to secondary radiation
  Scatter from within patient and collimator transmission are dominant sources near the treatment field
  Leakage through the linac head is the dominant source away from the treatment field
Discussed for IMRT, pediatrics,…
MLC leakage and transmission
Kragl et al ZMP 21 (2011)
See also Dzierma et al MP 39 (2012) for Siemens linac

## Slide 20: Peripheral dose
20
Number of segments and MUs comparable for FF and FFF beams
Dose reduction at 18 cm distance from the field size is ~ 15-20%
  Agreement between TLD-700, Farmer chamber & EBT films
Measurement Setup
Kragl et al ZMP 21 (2011)
IMRT Prostate case

## Slide 21: Radiation protection aspects – outside bunker
Measurements points in treatment room or around (different floors)
Floorplan Linac E:
23
18
22
Jank et al ZMP (2013)

## Slide 22: Some results …. (for MUW facility)
Floorplan Linac E:
23
18
22
Room door
Staircase
Linac console
Changing cubilce
Jank et al ZMP (2013)

## Slide 23: Summary & Conclusions
Convincing arguments of removing the FF for advanced treatment techniques (from C-arm linacs)
  dose rate, TPS modeling, peripheral dose, beam matching, …
Advanced RO techniques do not require flat dose profiles
  S(B)RT & IMRT are “prime applications”
State-of-the-art TPS can handle FFF beams
  Dose calculation accuracy comparable, if not better . . .
VMAT FF VMAT FFF

## Slide 24: Summary & Conclusions
Dedicated units without FF have been developed for “advanced“ RO
  Conventional C-arm linacs revisited & new developments will follow…
Efficient delivery gain only for hypo-fractionation
  Certainly new aspects w.r.t patient comfort and intra-fraction motion
Issues requiring follow up
  Standardization and QA, portal dosimetry, optimization, MLC design & leaf speed, radiobiology, small field dosimetry, …

## Slide 25
FFF team: Gabi Kragl, Wolfgang Lechner, Franziska Baier, Julia Jank, Tilo Wiezorek, ……

## Slide 26: Supplementary Material

## Slide 27: Treatment planning studies with FFF beams
True beam
  Scorsetti et al (2011): SBRT with RapidArc FFF, 70 patients, various tumor sites, showed feasibility of SBRT with FFF- beams w.r.t. acute toxicity
  Reggiori et al (2012): SBRT hepatic metastases, RapidArc 10MV FF and FFF, plans were dosimetrically equivalent, increase in MU but significant decrease in beam-on time for FFF
  Ong et al (2012): SBRT of lung and vertebra tumors, RapidArc 6MV FF and 10MV FFF, no major dosimetric differences, increase in MUs, average FFF beam delivery time 2.5 times lower than for FF- beams
  Nicoline et al (2012): VMAT with Rapid ARC FFF for advanced esophageal cancer, feasibility study, 20% reduced beam time compared to RA with FF
Tomotherapy
  Cozzi et al (2006) : comparison of IMRT, IMAT, HT, and Cyberknife for intracranial tumours, all techniques providing good OAR sparing, HT providing the best combination of target coverage indices
  Fogliata et al (2009): treatment of intracranial tumours with HT, IMRT and RapidArc, all three adequate to treat patients properly, RapidArc showed promising results, more studies necessary
  Lu et al (2012): dosimetric comparison of VMAT, HT and IMRT for nasopharyngeal carcinoma, similar PTV coverage and organ sparing for VMAT and HT but shorter delivery time for VMAT
27

## Slide 28: Treatment planning studies in literature
Prostate IMRT
  total MUs reduced by a factor of 2.0 for 6MV / 2.6 for 18MV
Fu et al 2004: absolute treatment time reduction depends on fraction dose
  beam on time reduced by 43%
Stathakis S et al. Radiat Isot (2009)
solid lines: flattened; dashed lines: unflattened
H&N IMRT
  Reduction from 590MU to 254MU
Vassiliev ON et al. IJROBP (2009)
solid lines: flattened; dashed lines: unflattened
Notes: Clinically acceptable plans can be created with existing TPS
Vassiliev - Eclipse
Stathakis - Pinnacle

## Slide 29: Treatment planning studies in literature
Prostate IMRT
  total MUs reduced by a factor of 2.0 for 6MV / 2.6 for 18MV
Fu et al 2004: absolute treatment time reduction depends on fraction dose
  beam on time reduced by 43%
Stathakis S et al. Radiat Isot (2009)
solid lines: flattened; dashed lines: unflattened
H&N IMRT
  Reduction from 590MU to 254MU
Vassiliev ON et al. IJROBP (2009)
solid lines: flattened; dashed lines: unflattened
Be aware of calibration effects:
1cGy / MU versus 2 - 3 cGy / MU
Notes: Clinically acceptable plans can be created with existing TPS
Vassiliev - Eclipse
Stathakis - Pinnacle

## Slide 30: ”Protective” Plate
Removes very low energies
Stabilises beam
Photon fluences Φγ, per initial electron, at the upstream surface of the water phantom as a function of photon energy Eγ. (FS 40x40 cm2).
Titt et al MP 33 (2006)
Electron fluences Φe, per initial electron, at the upstream surface of the water phantom as a function of electron energy Ee. (FS 40x40 cm2).
Dalary et al PMB 55 (2010)

## Slide 31: Dosimetric characteristics
Benchmark of FFF beams to „conventional“ photon beams
  FFF beams are still photon beams
Where is the difference ?
How does this difference influence dosimetry?

## Slide 32: SBRT delivery with FFF beams
For each image acquired ~ 2.5 / 5 Hz
  Generate two DRRs
  Compare it with the 2D pairs
  Iterate / Optimize
  Mean registration 5 (12 Hz)
WIP: validation of 6 DOF using kV + EPID images (SBRT pts)
Gendrin et. al. Med Phys (2011) / Furtado et al Acta Oncol (2013)
Clinical practice @ MedUni Wien: 15 Gy @ 65% isodose
    SBRT delivery time (same PTV coverage, similar dose to OARs)
  with FF - average time per beam ~ 40s
  without FF – average time per beam ~ 15s
+ EPID

## Slide 33: Data on intra-fraction motion prostate
Prostate and patients drift: before, during and after irradiation
Steiner et al IJROBP (2013)
Prostate
Pelvic bone

## Slide 34: Treatment planning @ MUW - WIP
Comparison of efficiency parameters:
  Median values of each patient and modality  Student’s t-test
Stochastic risk volume
 | IMRTFF | IMRTFFF | VMATFF | VMATFFF | diffIMRTFF-IMRTFFF | diffVMATFF-VMATFFF
Prostate |  |  |  |  |  | 
MU | 306 ± 29 | 344 ± 32 | 356 ± 41 | 386 ± 51 | +12%** | +8%**
T (s) | 405 ± 36 | 332 ± 37 | 143 ± 15 | 174 ± 16 | -18%** | +22*
V5Gy (%) | 44 ± 12 | 42 ± 12 | 46 ± 12 | 44 ± 12 | -5%** | -4%**
Head & Neck |  |  |  |  |  | 
MU | 654 ± 64 | 874 ± 105 | 738 ± 76 | 940 ± 124 | +34** | +27%**
T (s) | 723 ± 62 | 691 ± 64 | 301 ± 21 | 348 ± 21 | -4%* | +16%**
V5Gy (%) | 47 ± 10 | 46 ± 9 | 45 ± 9 | 45 ± 9 | -2% | 0%
FFF- MUs significantly increased compared to FF- MUs
IMRTFFF significantly faster than IMRTFF
VMATFF significantly faster than VMATFFF  MLC leaf speed restriction !
Lechner et al R&O (in press)

## Slide 35: Treatment planning @ MUW
Various studies concerning IMRTFF vs. VMATFF and Tomotherapy
Some studies on IMRT w/o FF
Little information for VMATFF vs. VMATFFF
No study considering all these techniques together
  Pareto front analysis
  Stochastic risk volume
  Delivery efficiency
Prostate / head-and-neck
VMAT FF VMAT FFF
IMRT FF IMRT FFF
  Localized prostate cases: 78Gy
  Head-and-neck cases: SIB 60Gy/50Gy

## Slide 36: SBRT and respiration managment
DIBH has the advantage of a high delivery efficiency but large number of MUs are “not practical”
Dose rate of FFF ~ 2 – 4 times FF dose rate
  Conventional beams with FF around 500 MU/min

## Slide 37: Photon beam spectra
FF FFF
FF FFF
Dalaryd, Kragl, Ceberg, Georg et al PMB 55 (2010)

## Slide 38: TPS commissioning – e.g. Monaco
Kragl et al R&O 100 (2011)

## Slide 39: Some more commissioning data . . .
39
Tomotherapy MC study
Jeraj et al (2004)
Cyberknife MC study
Araki et al (2006)
TrueBeam AAA
Chang et al (2013)

## Slide 40: Beam calibration & dose determination
Absolute dose measurements in solid water phantom following TRS 398 but with 0.5% kQ,Q0 correction for FFF beams
Setup for small field Dosimetry with SSD 95cmand d = 5cm, however for 10x10cm² reference FS (BrainLab m3),
 | 6MV FFF | 10 MV FFF
Farmer Chamber (Gy) | 16.42 | 17.46
Alanine (Gy) | 16.42 | 17.49
Difference (%) | 0.0 | -0.2

## Slide 41: Treatment planning @ MUW - WIP
Comparison of Pareto optimal fronts of FF and FFF:
  Pair wise comparison of FFF- plans to FF- plans (superior, inferior, comparable)
  Statistical analysis not straight forward
    Not normally distributed  non parametric test
    No adequate measure for differences between POFs  sign test
Stochastic risk volume
s/i/c | IMRTFF vs. IMRTFFF | VMATFF vs. VMATFFF | pIMRTFF-IMRTFFF | pVMATFF-VMATFFF
PR | 0/2/4 | 0/2/4 | ns | ns
HN | 3/0/3 | 5/0/1 | ns | ns
Lechner et al R&O (in press)

## Slide 42: Radiobiological experiments
In vitro radiobiological effects of high dose rates related to FFF beams
Challenge: homogeneous dose distribution for FF and FFF-beams
  Verify dose at cell layer/level
  EBT films frequently used @ MUW
42
γH2AX and DAPI
merged
2 cm

## Slide 43: Radiobiological experiments
In vitro radiobiological effects of high dose rates related to FFF beams
Challenge: homogeneous dose distribution for FF and FFF-beams
  Verify dose at cell layer/level
  EBT films frequently used @ MUW
43
γH2AX and DAPI
merged
2 cm

## Slide 44: SRT of uveal melanoma
>10 years experience and FU data
  LC and side effects comparable to proton therapy
  Report on > 200 patients (2 x Dunavölgyi et al IJROBP)
10 – 12 static mMLC shaped beams
Typically ~ 130 MU per beam
Increased dose rate with faster beam delivery less tiring for patient
Petersch et al MP 31 (2003)
mirror + scale
camera + housing
IR marker

## Slide 45: Beam calibration & dose determination
ABSORBED DOSE TO WATER FORMALISMS:
Limitations - small and non standard field
Andreo, PMB 45 (2000)
Xiong and Rogers MP 35 (2008)

## Slide 46: Material und Methoden
6 and 10 MV photon beams w/o FFF
  No neutron measurements
Points of interest in primary beam and scattered beam (23 POI)
FS 40 x 40 cm² and 15 x 15 cm²
Measurements w/o scatterer (30 x 30 x 30 cm² + 30 x 30 x 20 cm²)
Average results of measurments with both devices
Notes: 46

## Slide 47: Dose profiles
47
Penumbra values for flattened and unflattened beams agree within  1mm
Profile (off-axis dose) variation as a function of depth less pronounced for FFF beams
Less off-axis softening of beam
10 MV open FS [10 x 10cm²] with and without FF
Kragl et al R&O 93 (2009)
Notes: 47
First I want to present some dose profiles for the 15cm x 15cm field size for the 4 configurations, 6MV and 10MV and flattened and unflattened beams.
50% isodose:
Consider the 50% Isodose of the flattening filter free beam you can see, that it is smaller for 10MV than for 6MV. For the beam with flattening Filter the 50% Isodose is almost the same for 10MV and 6MV.
The 50% isodose is always larger for the flattened beam.
.
Shape:
Its obvious that the 10MV peak of the flattening filter free photon beam is more forward peaked than the 6MV beam. You see in this figure that the begin of the peak of the 6MV beam is at the point of 78% relative dose and for the 10MV beam it is at the point of 72% relative dose.
For the calculation of this point two tangents were cutted.

## Slide 48: Treatment planning @ MUW - WIP
Prostate cases

## Slide 49: Peripheral dose
49
Measurement Setup
Cashmore et al. IJROBP 80 (2011)
IMRT Pediatric intra cranial lesions
Plans clinically identical
Dose measurements in anthropomorphic phantom
Leakage dose up to 70% reduced

## Slide 50: Portal dosimetry
High dose rates of unflattened beams may lead to saturation
Softened energy spectrum of unflattened beams – more attenuation with increasing phantom thickness – increased response
Uniform attenuation across field – FFF beams more suitable for EPID dosimetry (smaller efforts for calibration necessary)
Winkler et al PMB 52 (2007)
Tyner et al. Med Phy 36 (2009)

## Slide 51: Peripheral dose
Cashmore et al IJROBP 80 (2011)
See also Kragl et al ZMP 21 (2011) for prostate IMRT, H&N IMRT, lung SBRT, …..

## Slide 52: Literature on Rad. Prot. effects for FFF is scarce ...
Vassiliev et al (Radiat. Prot. Dosim. 124(2) 2007)
  Varian Clinac® 21EX with 6 und 18 MV (2 mm Cooper)
    Dose ratio 6 MV (Primary beam): 2.0 ± 0.2 (MUW: data: 1.4 – 1.5)
    Dose ratio 6 MV (Scattered beam): 2.5 ± 0,2 (MUW data: 1.6 – 1.8)
    Dose ratio 18 MV (Primary beam): 3.5 ± 0.5 & less induced activity
    Dose ratio 18 MV (Scattered beam): 5.8 ± 0.8 & less induced activity
Aspects to consider:
    Treatment head design
    MU/cGy calibration
    Room size
    Measurement tools
Impact of FFF beams:
Higher allowed dose per week for existing treatment rooms
Less radiation to patients
Thinner walls for new rooms
Notes: 52

## Slide 53: Rotational therapy & intensity modulation
Dynamic arcs: subarcs with leafs adapting to target, constant collimator angle, constant doserate
IMAT: Rotational IMRT delivered on a conventional Linac with a conventional MLC
RapidArc: Varian`s IMAT solution that focuses on single arc delivery
VMAT: Elekta`s IMAT delivery solution
Cone Beam Therapy: Siemens work in progress
Smart Arc: Planning solution in Pinnacle
Hybrid Arc: Planning solution in iPlan
……….

## Slide 54: Spectral off axis changes
Lateral variation of the spectrum caused by combined effect of target and flattening filter
Tailor, R.C., et al. Med.Phys., 1998. 25(5): p. 662-667: Off axis HVL changes rather independent of linac and treatment head model
(for water)
Georg et al MP 37 (2010)
HVL measuring device
HVL measurement setup Nyström et al 1994, MM50 Microtron

## Slide 55: Reflection on increased dose rate
 | Elekta FFF | TrueBeam | Siemens FFF | Tomotherapy | Cyberknife
Max. dose rate MU /min | > 2000 | 2400 | 2000 | 850 | 600
Max field size | 40x40 cm² | 40x40 cm² | 40x40 cm² | 5x40 cm² | 6 cm
Available energies | 6/10 MV | 6/10 MV | 7 MV | 6 MV | 6 MV
55
Varian TrueBeam
Hrbacek et al. (2011)
Siemens Artiste
Dzierma et al. (2012)
Elekta Precise FFF

## Slide 56: Treatment head issues
Filter thickness optimization - trade-off between
  maximizing the dose rate and minimizing the scatter dose outside the field
  safety concerns and beam-steering properties
Dalaryd, Kragl, Ceberg, Georg et al PMB 55 (2010)
outer jaws
MLC
backup jaws
mirror
IC
primary collimator
FF or plate
target

## Slide 57: The ideal detector ….
No energy dependency
No dose rate dependency
No directional dependency
Small effective volume
Dose in more than 1 point
Online display
High sensitivity
………….
Ideality
Reality

## Slide 58: Dose rate dependence of detectors
OF measured with dose rate corrected diamond detector vs. alanine
58
After correction: output factors measured with diamond detector and alanine agree within 2%
Lechner et al R&O 109 (2013)
Notes: 58
Typical effects of changing the dose rate within every dose pulse (SSD).
The diode actually gives a higher signal for higher dose rates (so-called super linearity?). Results differ...
LIC not intended for pulsed radiation, i.e. not a fair test.

## Slide 59: Remarks on small field dosimetry
Advanced photon beam delivery methods based on small fields
  Beam calibration conditions can be different
Practical pitfalls…
Extension of existing protocols
  Upcoming IAEA TECDOC
Experience from IMRT audits
  RPC and others
Adapted from Das et al. (2009) Med Phys 35:206-215
Alfonso et al. (2008) Med Phys 35:5179-5186

## Slide 60: Small Field Dosimetry – Volume averaging effects
Volume averaging correction factors ~ few %
Example of 3 detectors with different volumes:
  PFD , PinPoint16 and CC13
  More pronounced for detectors with larger volume
  Negligible for field sizes > 12x12mm²
Lechner et al R&O 109 (2013)

## Slide 61: Recombination effects of FFF beams
Two voltage method applicable for FFF-beams?
  Agreement of two voltage method (Jaffé-plot: 1/V vs 1/Q) within 0.3%
Liquid filled ionization chamber for FFF-beams?
  Collection efficiency above 0.99 for FF-beams but decrease to 0.94 for FFF-beams
  Might not be suitable for FFF-beams
Kry et al. (2012) J Appl Clin Med Phys 13:318- 325
Lang et al. (2012) PMB 57:2819-2827

## Slide 62: REMINDER - Jaffé plot
Chart: 1/V
| 0.0025 | 0.00333333333333334 | 0.00500000000000001 | 0.00666666666666667 | 0.01 | 0.02 | 0.0333333333333333 | 0.0666666666666667
Mean(M) | 3.1437 | 3.1538 | 3.1671 | 3.1703 | 3.1765 | 3.1881 | 3.201 | 3.2343
 | 3.1671 | 3.1703 | 3.1765 | 3.1881 | 3.201 | 3.2343
Charge multiplication effect at 300V and 400V

## Slide 63: Measurement of surface doses
FFF is scatter source and source of contaminating electrons
Surface dose measurements require high spatial resolution
  films, extrapolation chamber
Differences occur at very shallow depth
Dose prediction at shallow depth by TPS is (mostly) weakness
De Puysseleyr, Lechner (unpublished)

## Slide 64: Radiation Protection aspects – inside bunker
Treatment head leakage radiation
Cashmore PMB 53 (2008)
Kragl ZMP 21 (2011)

## Slide 65: Room shielding - % reductions (for MUW facility)
ÖNORM S 5216 „Medizinische Elektronen-beschleuniger-Anlagen - Strahlenschutzregeln für die Errichtung“
NCRP Report No. 151 „Structural Shielding Design and Evaluation for Megavoltage X- and Gamma-Ray Radiotherapy Facilities“
Nominal Energies
  6.2  5.6 MV
  10.5  8.5 MV
Leakage reduction:
  – 52 % @ 6MV
  – 65 % @ 10 MV
Jank et al ZMP (2013)

## Slide 66: Penumbra definition and comparison
Renormalization of profiles to inflection points
66
Pönisch et al. MP 33 (2006)
Renormalization of profiles to maximum of the third derivative
Fogliata et al. MP 39 (2012)
penumbra FFF ~ penumbra FF, i.e. difference < 1 mm
Notes: for determination of field size and penumbras
renormalization